Rigaku D/Max – B x-ray diffractometer instrument specs (general)

Make:  Rigaku

Model: D/Max - B

Copper target X-ray tube (40 KV and 30 mA)

Copper Kradiation  (composed of Cu K1 = 1.5405 Å and Cu K2 = 1.5443 Å; the weighted average of the two is Cu Kavg = 1.54184 Å)

Bragg Brentano parafocusing geometry

Diffracted beam monochromator (curved graphite crystal)

monochromator angle = 13.27 degrees (2theta = 26.57 degrees)

1 degree divergence slit

1 degree scatter slit

.3 mm receiving slit

.8 mm monochromator receiving slit

Incident and diffracted beam Soller slits (angular aperture = 5 degrees)

goniometer radius = 185mm

Example:

“X-Ray diffraction data were taken with a Rigaku D/Max-B x-ray diffractometer with Bragg – Brentano parafocusing geometry, a diffracted beam monochromator, and a conventional copper target x-ray tube set to 40 KV and 30 mA.”

Rigaku D/Max-B Instrument parameters (detailed)

X-ray tube focus (normal)								0.1 x 10 mm

Take-off angle										6˚

Goniometer radius (source to sample =  sample to RS)					185 mm

Sample to SS										141 mm

DS to sample										~90 mm

DS height										10 mm

DS width (typical)									1˚

RS height										20 mm

RS width (typical)									0.3˚

Soller slit (incident beam) angular aperture						5˚

l~ 10mm, s ~ .5 mm (2 = 2tan-1 (s/l) = 5.72˚)

Soller slit (diffracted beam) angular aperture						5˚

l = 25 mm, s = 1mm  (2 = 2tan-1 (s/l) = 4.58˚)

Effective diffracted beam angular aperture w/out soller slit(2)				7.94˚

s = goniometer radius + monochromator beam path = 288 mm

l = RSM height = 20 mm

Effective diffracted beam angular aperture w/out mono and soller slit(2 		12.3˚

s = goniometer radius = 185mm, l = RS height = 20 mm

Fixed sample irradiation height								~16 mm

Diffracted-Beam monochromator

Curved graphite crystal with [0002]  d = 6.708 Å; curvature = 224 mm

RS to crystal  distance									51.5 mm

RSM to crystal distance									51.5 mm

											13.27˚

											26.57˚

Items in red are educated guesses or experimental measurements with some degree of uncertainty